## Slide 1: Fundamental Techniques and Measurements
Mass Measurements
Volume Measurements
Preparation of a solution of known concentration
UV-Visible Spectrophotometer

## Slide 2: Electronic Balance
600 g
What does an electronic balance measure? _____
If you took an electronic balance with a capacity of 100 g to the moon what would its range be? _____
force
Why do good electronic balances have bubble level indicators?

## Slide 3: Mass: Electronic Balance
Accuracy
  4 to 6 significant digits
Calibration
  Use known mass
  Check weekly or when balance is moved
gravity
Sources of error
  Balance must be calibrated and maintained in same orientation in _________ field
  hydroscopic chemicals: dry to constant mass first (will increase in mass rapidly as they reabsorb water on the balance!)
  When preparing a solution of a given concentration it may be difficult to get the exact mass desired____________
  evaporation of wet samples
report actual!

## Slide 4: Electronic Balance
For maximum accuracy use balance with _______ resolution possible!
Don’t forget to clean the balance if you spill any chemicals!!!!!!
M
o
d
e
l
C
a
p
a
c
i
t
y
R
e
s
o
l
u
t
i
o
n
D
I
-
1
0
0
1
0
0
g
0
.
0
0
0
1
g
D
I
-
8
0
0
8
0
0
g
0
.
0
1
g
D
I
-
5
0
0
0
5
0
0
0
g
0
.
1
g
smallest

## Slide 5: Volume
Volumetric flask
  accuracy of ______/100 mL
Graduated cylinder
  accuracy of ______/100 mL
Beaker
  accuracy of _____/100 mL
Pipette
  accuracy of ± _____ for 100-1000 µL
  accuracy of ± _____ for 10-100 µl
0.16 mL
1 mL
5 mL
0.6%
0.8%
What will accuracy of solution be if you use pipette, volumetric flask, and electronic balance? ________
1 %
What controls the accuracy? _______
Pipette

## Slide 6: Digital Pipettes
Air displacement
Do not directly contact fluid volume
    avoids contamination of pipette
    avoids sample carryover
Require air tight connection between tip and body

## Slide 7: Pipette Workings
piston
cylinder
Pipette tip

## Slide 8: Preparation of Solutions
Example: Prepare 100 mL of a 30 mM solution of methylene blue.
The molecular weight of methylene blue (C16H18N3SCl) is 319.87 g.
concentration
volume
conversion
=
mass

## Slide 9: Preparation of Solutions
Fill volumetric flask half way with distilled water
Add reagent (could be solid or liquid)
Mix
Fill volumetric flask to the line
Mix
Verify that volume didn’t change (if necessary refill to line)

## Slide 10: UV-Visible Spectrophotometer
Theory
Instrument
Sample requirements
Software

## Slide 11: Light Attenuation by an Aqueous Solution
P0
P
P0
P
P is light intensity (photons/s)

## Slide 12: Theory: Light Attenuation = f(?)
For a given excitation process, a molecule absorbs only one discrete amount of energy: expect very narrow absorption lines.
Different vibrational and rotational states yield _______ absorption lines.
First order decay with distance
broad

## Slide 13: A=bc
Po - _________ light intensity
P light intensity after passing through sample
b - ______________
c - ______________
 - ___________ coefficient (function of wavelength and molecule)
incident
path length (1 cm)
concentration
extinction
e is not a function of concentration!

## Slide 14: Absorption Spectra
Absorption Spectra for Methylene Blue
  Broad peaks
  Absorbs ________, looks ______
blue
orange

## Slide 15: Sample Requirements
Sipper cell
  peristaltic pump draws sample into sipper cell
  requires a few mL to displace previous cell contents
Light source
detector
pump
sample

## Slide 16: Absorbance Measurement Limitations
Po is a function of the _____.
If absorbance is high what is P? ______
Suppose A = 3, what is Po/P? _____
Suppose I create samples of higher and higher concentration. What will happen to the absorbance measurements?
lamp
small
1000
There is a _________ (non zero) P that can be measured by an instrument.
minimum
A _______ keep increasing!
doesn’t
Amax 3